- Pressure Transducer Error (full scale voltage of 0.5 for both):

  - Minimum resolution error

    - Static pressure

      - Full scale pressure: 60 psi

      - Full scale voltage: 0.5V

      - Bits: 12

      - Steps: 2^Bits = 4096

      - Step size: Steps / Full scale voltage = 1e-04

      - Minimum resolution: Slope * Step size = 0.0183 psi

      - Minimum resolution = 0.3947 kPa

    - Stagnation pressure

      - Full scale pressure: 300 psi

      - Full scale voltage: 0.5V

      - Bits: 12

      - Steps: 2^Bits = 4096

      - Step size: Steps / Full scale voltage = 1e-04

      - Minimum resolution: Slope * Step size = 0.0732 psi

      - Minimum resolution = 1.579 kPa

  - Linearity

    - Static: 60 psi * 0.0025 = 0.15 psi

    - Stagnation: 300 psi * 0.0025 = 0.75 psi

  - Repeatability

    - Static: 60 psi * 0.002 = 0.12 psi

    - Stagnation: 300 psi * 0.002 = 0.6 psi

  - Temperature Shift

    - Static: 60 psi * 0.0075 = 0.45 psi

    - Stagnation: 300 psi * 0.075 = 2.25 psi

  - Null Shift

    - Static: 60 psi * 0.005 = 0.3 psi

    - Stagnation: 300 psi * 0.005 = 1.5 psi

  - Total

    - SQRT(Minimum resolution^2 + Linearity^2 + Repeatability^2 + Temperature Shift^2 + Null Shift^2)

    - Static: +/-0.81 psi

    - Stagnation: +/-4.06 psi